Управление частотами MCR на низком уровне, используя команды SCPI.

Нам требуется настроить следующие параметры системы:

|  | Переменная | Тип | Описание |
| --- | --- | --- | --- |
| 1 | FLO | float | Частота выходного сигнала ГУН |
| 2 | FSyn | float | Частота опорного генератора |

Приемник AVDN-01M-L1 реализует частоту ГУНа в диапазоне 1550-1700 МГц, для AVDN-01M-L2 частоты ГУНа 1100-1300 МГц.

Гетеродин приемника реализован  на основе синтезатора частот с дробно-переменным делителем частоты (Fractional-N) HMC704LP4E.

FLO=FSynRNint+Nfrac224, где FSyn=FTW248⋅1000MHz.

FTW – значение, управляющее частотой опорного генератора, диапазон [0…248-1].

R – значение делителя частоты опорного генератора; если FSyn>96МГц, то R=3, иначе R=2.

Nint – значение целой части петлевого делителя частоты, диапазон [20…219-5].

Nfrac – значение дробной части петлевого делителя частоты, диапазон [0…224-1].

Команды SCPI для управления вышеобозначенными параметрами:

SYNthesizer:FTW

Команда для загрузки значения FTW.

SYNthesizer:FTW?

Команда для считывания текущего значения FTW.

SYNthesizer:PTW

Команда для загрузки значения PTW. Это значение управляет фазой сигнала с опорного генератора. Диапазон значений [0…214-1], соответсвует 360 градусам.

SYNthesizer:PTW?

Команда для считывания текущего значения PTW.

DEMn:RDIV

Команда для загрузки значения R n-го демодулятора.

DEMn:RDIV?

Команда для считывания текущего значения R n-го демодулятора.

DEMn:NINT

Команда для загрузки значения Nint n-го демодулятора.

DEMn:NINT?

Команда для считывания текущего значения Nint n-го демодулятора.

DEMn:NFRAC

Команда для загрузки значения Nfrac n-го демодулятора.

DEMn:NFRAC?

Команда для считывания текущего значения Nfrac n-го демодулятора.

Рекомендуется сначала прогрузить систему, используя высокоуровневые команды SYNthesizer:FREQuency и DEMn:FREQuency, а затем подкорректировать их, используя вышеобозначенные команды.

При считывании значения частоты командами SYNthesizer:FREQuency? и DEMn:FREQuency?, выводится прежнее значение частоты, которое было загружено этими же командами.

3 Firmware Update

MCU which controls the synthesizer RF-module has two types of nonvolatile memory: Flash that contains program code, and EEPROM which stores all information about the device, including information about current firmware version. Thus the firmware update consists of two files in Intel Hex format, one – for Flash, and the other – for EEPROM. Firmware update is carried out through RS-232 or USB interface, connected to the PC. The connection via USB is implemented with aid of USB to UART bridge, so from the view point of PC software it looks like a COM-port, fig. 5. This method requires driver installation 2 (integral circuit CP2102). After driver installation and when the instrument is connected to the PC, the COM-port appears in the hardware list. Its number is indicated at the end of the string (fig. 5).

Figure 5: Device Manager window (My Computer .Manage)

When powering up the evaluation board, it enters firmware update mode and waits for the appropriate command from the PC. If it doesn’t receive the command within a half of a second, it exits the update mode and no longer responds to the read/write Flash or EEPROM commands, since in regular mode RS-232 and USB interfaces are used for remote control. For the firmware updating the XMI Programmer (XMEGA Instrument Programmer) application

is used, fig. 6. The sequence of actions is the following:

Figure 6: Firmware update application – XMI Programmer

- Turn off the instrument;

- Connect it to the PC via RS-232 or USB cable;

- Launch the XMI Programmer software;

- Select the port, then press Connect button. If you chose the wrong port, then press the Stop button, select the right port, then press Connect;

- Turn on the instrument. The connection should be established within a second (you will see the string Connecting... Ok!). If it’s not happened then probably wrong port was selected. In such a case press the Stop button, turn off the instrument, change the port in drop-down list, press Connect and then turn on the instrument. The string Connecting... Ok! should be displayed;

- Select the Write to MCU radio-box, the Rewrite All EEPROM Data flag should remain inactive; If you select this flag, then all EEPROM memory will be re written, not only the part, that contains the information about current firmware revision. That is the personal data of the instrument (serial number, operation time, power-on count) will be erased, which is not desirable

- Select the files for Flash and EEPROM;

- Press Go! button. The firmware update process takes about 30 seconds and comprises the data integrity check, the results of the operations are displayed in the log window;

- After successful verifying the written data you may close the application. You can save the current firmware and EEPROM data by selecting Read from MCU radio-button at 6-th step.